## GR&R study
Gage R&R study — crossed, 10 parts × 3 appraisers × 3 trials
Fill the yellow cells: the study details, the specification limits and all 90 readings. Number the parts underneath, randomise the order for every trial, and keep the appraisers blind to each other's (and their own earlier) readings.
Gauge / asset no. | [7 empty] | Lower spec limit (LSL)
Characteristic | [7 empty] | Upper spec limit (USL)
Part number | [7 empty] | Process std dev (optional) | [5 empty] | ← long-term σ, if known
Study date / by | [7 empty] | Appraisers A / B / C
Tolerance (USL − LSL)
Readings
Part | Appraiser A |  |  | Appraiser B |  |  | Appraiser C |  |  | Calculation columns (do not edit)
no. | Trial 1 | Trial 2 | Trial 3 | Trial 1 | Trial 2 | Trial 3 | Trial 1 | Trial 2 | Trial 3 | Part mean | Mean A | Mean B | Mean C | Range A | Range B | Range C
1
2
3
4
5
6
7
8
9
10
Appraiser mean | [9 empty] | Avg range
Readings entered | Appraiser A Trial 3: 0; Appraiser B Trial 1: Enter all 90 readings to see results
ANOVA method (preferred) | Appraiser C Trial 3: Average & range method
Source |  | DF | SS | MS | F | p-value |  |  | Quantity |  | Value | % tolerance | % study var.
Parts | Appraiser A Trial 2: 9; Appraiser C Trial 3: Average range R̿
Appraisers | Appraiser A Trial 2: 2; Appraiser C Trial 3: Appraiser-mean spread X̄diff
Part × appraiser | Appraiser A Trial 2: 18; Appraiser C Trial 3: Part-mean range Rp
Repeatability | Appraiser A Trial 2: 60; Appraiser C Trial 3: EV = R̿ × K1
Total | Appraiser A Trial 2: 89; Appraiser C Trial 3: AV = √((X̄diff × K2)² − EV²/(n·r))
GRR = √(EV² + AV²)
Interaction | Appraiser C Trial 3: PV = Rp × K3
TV = √(GRR² + PV²)
Component |  | Variance | Std dev | 6 × SD | % tolerance | % study var. | % process |  | ndc (truncated)
Repeatability (EV) | Appraiser C Trial 3: K1 (3 trials) |  | 0.5908
Reproducibility (AV) | Appraiser C Trial 3: K2 (3 appraisers) |  | 0.5231
of which interaction | Appraiser C Trial 3: K3 (10 parts) |  | 0.3146
Gage R&R (GRR) | Appraiser C Trial 3: K constants are the published values for this 10 × 3 × 3 design (6σ basis). The A&R method cannot detect a part × appraiser interaction — when ANOVA flags one, trust ANOVA.
Part variation (PV)
Total variation (TV)
ndc = 1.41 × PV ÷ GRR, truncated
Verdict (ANOVA)
GRR vs tolerance (use for accept / reject decisions)
GRR vs study variation (process control — parts must span the process)
GRR vs process std dev (if entered)
Number of distinct categories
Where the error lives
Decision / actions / customer agreement | bar: Appraiser C Trial 2 to Appraiser C Trial 3
Qlause · qlause.app · Interpretive guidance only — follow the AIAG MSA manual / VDA Volume 5 and your customer-specific requirements; work from your licensed copies.

## GR&R worked example
Gage R&R study — crossed, 10 parts × 3 appraisers × 3 trials (worked example)
Worked example from the Qlause article 'MSA and Gage R&R explained': height gauge HG-003 on the pump mounting bracket's formed height, 24.5 ± 0.30 mm. Every result below is a live formula — change a reading and watch it move.
Gauge / asset no. |  | HG-003 height gauge, 0.01 mm resolution | [5 empty] | Lower spec limit (LSL) |  |  | 24.2
Characteristic |  | Formed height (OP10) | [5 empty] | Upper spec limit (USL) |  |  | 24.8
Part number |  | 12345 / Rev C — pump mounting bracket | [5 empty] | Process std dev (optional) | [5 empty] | ← long-term σ, if known
Study date / by |  | Worked example | [5 empty] | Appraisers A / B / C |  |  | Setter 1 / Setter 2 / Setter 3
[8 empty] | Tolerance (USL − LSL) |  |  | 0.6
Readings
Part | Appraiser A |  |  | Appraiser B |  |  | Appraiser C |  |  | Calculation columns (do not edit)
no. | Trial 1 | Trial 2 | Trial 3 | Trial 1 | Trial 2 | Trial 3 | Trial 1 | Trial 2 | Trial 3 | Part mean | Mean A | Mean B | Mean C | Range A | Range B | Range C
1 | 24.38 | 24.39 | 24.42 | 24.41 | 24.38 | 24.4 | 24.37 | 24.39 | 24.36 | 24.3889 | 24.3967 | 24.3967 | 24.3733 | 0.04 | 0.03 | 0.03
2 | 24.56 | 24.56 | 24.59 | 24.58 | 24.58 | 24.55 | 24.59 | 24.52 | 24.57 | 24.5667 | 24.57 | 24.57 | 24.56 | 0.03 | 0.03 | 0.07
3 | 24.47 | 24.47 | 24.47 | 24.48 | 24.5 | 24.49 | 24.5 | 24.44 | 24.47 | 24.4767 | 24.47 | 24.49 | 24.47 | 0 | 0.02 | 0.06
4 | 24.62 | 24.64 | 24.6 | 24.64 | 24.65 | 24.62 | 24.64 | 24.63 | 24.64 | 24.6311 | 24.62 | 24.6367 | 24.6367 | 0.04 | 0.03 | 0.01
5 | 24.46 | 24.42 | 24.43 | 24.45 | 24.47 | 24.47 | 24.42 | 24.42 | 24.42 | 24.44 | 24.4367 | 24.4633 | 24.42 | 0.04 | 0.02 | 0
6 | 24.51 | 24.52 | 24.52 | 24.57 | 24.53 | 24.53 | 24.52 | 24.5 | 24.51 | 24.5233 | 24.5167 | 24.5433 | 24.51 | 0.01 | 0.04 | 0.02
7 | 24.59 | 24.6 | 24.57 | 24.63 | 24.6 | 24.61 | 24.6 | 24.59 | 24.6 | 24.5989 | 24.5867 | 24.6133 | 24.5967 | 0.03 | 0.03 | 0.01
8 | 24.37 | 24.37 | 24.38 | 24.37 | 24.4 | 24.37 | 24.34 | 24.4 | 24.36 | 24.3733 | 24.3733 | 24.38 | 24.3667 | 0.01 | 0.03 | 0.06
9 | 24.52 | 24.57 | 24.53 | 24.57 | 24.54 | 24.55 | 24.51 | 24.5 | 24.52 | 24.5344 | 24.54 | 24.5533 | 24.51 | 0.05 | 0.03 | 0.02
10 | 24.47 | 24.48 | 24.51 | 24.47 | 24.5 | 24.47 | 24.46 | 24.44 | 24.44 | 24.4711 | 24.4867 | 24.48 | 24.4467 | 0.04 | 0.03 | 0.02
Appraiser mean | Appraiser A Trial 1: 24.4997; Appraiser B Trial 1: 24.5127; Appraiser C Trial 1: 24.489 | Avg range |  |  |  | 0.029 | 0.029 | 0.03
Readings entered | Appraiser A Trial 3: 90; Appraiser B Trial 1: Complete — results below are live
ANOVA method (preferred) | Appraiser C Trial 3: Average & range method
Source |  | DF | SS | MS | F | p-value |  |  | Quantity |  | Value | % tolerance | % study var.
Parts | Appraiser A Trial 2: 9; Appraiser A Trial 3: 0.5986; Appraiser B Trial 1: 0.066512; Appraiser C Trial 3: Average range R̿ |  | 0.029333
Appraisers | Appraiser A Trial 2: 2; Appraiser A Trial 3: 0.0084289; Appraiser B Trial 1: 0.0042144; Appraiser C Trial 3: Appraiser-mean spread X̄diff |  | 0.023667
Part × appraiser |  | 18 | 0.0060156 | 0.0003342 | 1.0704 | 0.40237 |  |  | Part-mean range Rp |  | 0.25778
Repeatability | Appraiser A Trial 2: 60; Appraiser A Trial 3: 0.018733; Appraiser B Trial 1: 0.00031222; Appraiser C Trial 3: EV = R̿ × K1 |  | 0.01733 | 17.33% | 20.68%
Total | Appraiser A Trial 2: 89; Appraiser A Trial 3: 0.63178; Appraiser C Trial 3: AV = √((X̄diff × K2)² − EV²/(n·r)) |  | 0.011969 | 11.97% | 14.28%
| Appraiser C Trial 3: GRR = √(EV² + AV²) |  | 0.021062 | 21.06% | 25.14%
Interaction | Appraiser A Trial 2: Pooled (p > 0.25); Appraiser C Trial 3: PV = Rp × K3 |  | 0.081097 |  | 96.79%
| Appraiser C Trial 3: TV = √(GRR² + PV²) |  | 0.083787
Component |  | Variance | Std dev | 6 × SD | % tolerance | % study var. | % process |  | ndc (truncated) |  | 5
Repeatability (EV) |  | 0.00031729 | 0.017813 | 0.10688 | 17.81% | 20.17% |  |  | K1 (3 trials) |  | 0.5908
Reproducibility (AV) |  | 0.00012991 | 0.011398 | 0.068386 | 11.4% | 12.9% |  |  | K2 (3 appraisers) |  | 0.5231
of which interaction |  | 0 | 0 | 0 | 0% | 0% |  |  | K3 (10 parts) |  | 0.3146
Gage R&R (GRR) |  | 0.0004472 | 0.021147 | 0.12688 | 21.15% | 23.94% |  |  | K constants are the published values for this 10 × 3 × 3 design (6σ basis). The A&R method cannot detect a part × appraiser interaction — when ANOVA flags one, trust ANOVA.
Part variation (PV) | Appraiser A Trial 2: 0.0073549; Appraiser A Trial 3: 0.085761; Appraiser B Trial 1: 0.51457; Appraiser B Trial 3: 97.09%
Total variation (TV) | Appraiser A Trial 2: 0.0078021; Appraiser A Trial 3: 0.08833; Appraiser B Trial 1: 0.52998; Appraiser B Trial 3: 100%
ndc = 1.41 × PV ÷ GRR, truncated | Appraiser B Trial 2: 5
Verdict (ANOVA)
GRR vs tolerance (use for accept / reject decisions) | Appraiser C Trial 1: 21.15%; Appraiser C Trial 2: 10–30% — may be acceptable: record the decision and reasons
GRR vs study variation (process control — parts must span the process) | Appraiser C Trial 1: 23.94%; Appraiser C Trial 2: 10–30% — may be acceptable: record the decision and reasons
GRR vs process std dev (if entered)
Number of distinct categories | Appraiser C Trial 1: 5; Appraiser C Trial 2: ndc ≥ 5 — adequate discrimination for process control
Where the error lives | Appraiser C Trial 2: Repeatability dominates — look at the instrument, fixturing and how the part sits / where it is measured
Decision / actions / customer agreement | Appraiser C Trial 2: Marginal (21% of tolerance, ndc 5). Repeatability-dominated; type 1 on the master is 3× tighter, so the scatter is part seating, not the instrument. Action: locating nest + defined measuring point on the WI; re-study after the change. Accepted for use meanwhile with the X-bar/R chart.
Qlause · qlause.app · Interpretive guidance only — follow the AIAG MSA manual / VDA Volume 5 and your customer-specific requirements; work from your licensed copies.

## Type 1 study
Type 1 gauge study
One appraiser, one reference part of known value, the same feature measured repeatedly (50 readings is common practice, 25 a minimum), lifting and replacing the part each time. Fill the yellow cells.
Gauge / asset no.
Reference value
LSL
USL
Reading | +1 | +2 | +3 | +4 | +5 | Result | Value | Interpretation
1–5 | [5 empty] | Readings (n) | 0
6–10 | [5 empty] | Tolerance T
11–15 | [5 empty] | Mean
16–20 | [5 empty] | Standard deviation s
21–25 | [5 empty] | Bias (mean − reference)
26–30 | [5 empty] | Bias as % of tolerance
31–35 | [5 empty] | Bias t-statistic
36–40 | [5 empty] | Bias p-value (two-sided)
41–45 | [5 empty] | Cg = 0.2·T ÷ 6s
46–50 | [5 empty] | Cgk = (0.1·T − ¦bias¦) ÷ 3s
Cg / Cgk use the common 20% / 10% of tolerance convention with a 1.33 target; your customer may set others. A type 1 study checks the instrument on a master — it does not replace a Gage R&R on real parts with the real appraisers. Keep resolution at or below a tenth of the tolerance.
Qlause · qlause.app · Interpretive guidance only — follow the AIAG MSA manual / VDA Volume 5 and your customer-specific requirements; work from your licensed copies.

## Type 1 example
Type 1 gauge study (worked example)
HG-003 on a certified reference part (24.512 mm): one setter, 50 readings, part lifted and replaced each time.
Gauge / asset no. |  | HG-003
Reference value |  | 24.512
LSL |  | 24.2
USL |  | 24.8
Reading | +1 | +2 | +3 | +4 | +5 | Result | Value | Interpretation
1–5 | 24.51 | 24.51 | 24.51 | 24.52 | 24.52 | Readings (n) | 50
6–10 | 24.51 | 24.52 | 24.52 | 24.52 | 24.52 | Tolerance T | 0.6
11–15 | 24.51 | 24.51 | 24.52 | 24.52 | 24.52 | Mean | 24.515
16–20 | 24.51 | 24.52 | 24.52 | 24.51 | 24.51 | Standard deviation s | 0.0058029
21–25 | 24.53 | 24.51 | 24.51 | 24.52 | 24.51 | Bias (mean − reference) | 0.003
26–30 | 24.51 | 24.51 | 24.51 | 24.5 | 24.52 | Bias as % of tolerance | 0.5%
31–35 | 24.52 | 24.52 | 24.51 | 24.52 | 24.52 | Bias t-statistic | 3.6556
36–40 | 24.52 | 24.52 | 24.51 | 24.51 | 24.51 | Bias p-value (two-sided) | 0.0006252 | Bias statistically significant — judge its size against the tolerance
41–45 | 24.52 | 24.52 | 24.52 | 24.52 | 24.51 | Cg = 0.2·T ÷ 6s | 3.4466 | ≥ 1.33 — capable
46–50 | 24.51 | 24.52 | 24.51 | 24.51 | 24.51 | Cgk = (0.1·T − ¦bias¦) ÷ 3s | 3.2742 | ≥ 1.33 — capable
Cg / Cgk use the common 20% / 10% of tolerance convention with a 1.33 target; your customer may set others. A type 1 study checks the instrument on a master — it does not replace a Gage R&R on real parts with the real appraisers. Keep resolution at or below a tenth of the tolerance.
Qlause · qlause.app · Interpretive guidance only — follow the AIAG MSA manual / VDA Volume 5 and your customer-specific requirements; work from your licensed copies.

## How to use
How to use this template
1 | Scope the programme from your control plans: every gauge or test system named in an evaluation-method column needs a study, starting with those on special characteristics.
2 | Check resolution first: the gauge should read to a tenth of the tolerance (or process spread) or finer.
3 | Run a Type 1 study on a reference part. If Cg or Cgk is below 1.33, fix the instrument before spending time on a full study.
4 | Choose ten parts that span the real process range — from a week of production, not ten consecutive parts. Parts that are too alike make any gauge look bad against study variation; hand-picked extremes make any gauge look good.
5 | Use the people who measure in production, with the production method, in the production location. Number parts underneath, randomise the order for every trial, record the readings for them.
6 | Enter all 90 readings on the 'GR&R study' sheet. ANOVA is the method to report; the average & range block is there for comparison and hand-checking.
7 | Pick the right denominator: % tolerance when the gauge accepts or rejects parts; % study variation (or % of a known process standard deviation) when it feeds SPC, capability or set-up decisions.
8 | Read the bands: under 10% generally acceptable; 10–30% may be acceptable depending on the application, cost and customer agreement — record the decision; over 30% improve the system. ndc should be 5 or more.
9 | Use the EV / AV split to find the cause: repeatability points at the instrument, fixturing and part seating; reproducibility at method and technique.
10 | Re-study after events: a new or repaired gauge, a new fixture or method, a tolerance or process change, an out-of-control stability chart, or a complaint on a characteristic the gauge checks.
Attribute (go/no-go) gauges need an attribute agreement study instead — effectiveness, miss rate, false alarm rate and kappa against reference decisions. The article explains how.
Yellow cells are inputs (blue text = example values). Everything else is a formula — protect the sheet if you circulate it.
Read the full guide: qlause.app/consulting/insights/msa-gage-rr-explained · Control Plan Template, Process Flow Template and the PFMEA Toolkit: qlause.app/templates
Qlause · qlause.app · Interpretive guidance only — follow the AIAG MSA manual / VDA Volume 5 and your customer-specific requirements; work from your licensed copies.